PURPOSE:

To ensure scales are calibrated properly.

SCOPE:

This process involves repack and mix plant operators.

DOCUMENTS:

Form 4, calibration check – scale.

PROCEDURE:

On the 1st Monday of every month, check the accuracy of the scales with the test weights and record the results in the scale log book on form 4.

- Obtain Calibration Weights and log book from Repack cupboard.

- Ensure scale is empty and level.

- Zero the scale.

- Place lowest weighted of the calibration weights for that specific scale on to plate, read and record displayed weight.

- Take highest weighted of the calibration weights for that specific scale and place on plate, read and record displayed weight.

- If recorded weights are plus or minus 1% of the calibration weight, i.e. 1g per 1kg, then it is acceptable.  Otherwise advise deviation to operations and quality control to determine if scale is to be in continued use.

- For scales greater than 10kg it is advisable to use 3 weights to calibrate, low, medium and high.

- Annually – when the outside consultant conducts the annual verification and calibration of scales, have them verify a 15kg and 20kg test weight.

RESPONSIBILITIES:

Mix plant operator and repack staff to do check and keep record once per month.

Operations Manager in the event of a fail - to arrange for scales to be serviced.

REVIEW AND RECORDS:

As needed